## Slide 1
Cryocourse 2016
School and Workshop in Cryogenics and Quantum Engineering
26th September - 3rd October 2016
Aalto University, Espoo, Finland
Introduction to Dilution Refrigerators
Part I-Thermodynamics
Henri GODFRIN
CNRS/IN/MCBT – Grenoble

## Slide 2: This part of the course is based on:
Basic operation of cryocoolers and related thermal machines
A.T.A.M. de Waele
J. of Low Temp. Physics, Vol.164, pp.179-236 (2011) (open access)
Introduction to dilution refrigeration
A.T.A.M. de Waele
Lectures given at Cryocourse 2013 and former ones
Dilution refrigerators
H. Godfrin
Lectures given at Cryocourse 2013 and former ones
Experimental Principles and Methods Below 1 K
Lounasmaa, O.V.
Academic Press, London, p. 316 (1974).
Matter and Methods at Low Temperatures
Pobell, Frank
Springer-Verlag, Berlin, p. 461 (2007).
Wikipedia: https://en.wikipedia.org/wiki/Dilution_refrigerator

## Slide 3: Outline
Introduction
Schematic diagram
Thermodynamics of dilution refrigeration
Thermal properties of pure 3He
Properties of 3He-4He mixtures
The 1 K bath
Why a still heater
The mixing chamber
Osmotic enthalpy
Limiting temperature of dilution refrigerators

## Slide 4: History
The dilution refrigerator principle was suggested by Heinz London in 1952
H. London, G.R. Clarke, and E. Mendoza proposed a prototype of continuous refrigerator in 1962
It was realized in 1964 in the Kamerlingh Onnes Laboratorium at Leiden University
Das, P.; Ouboter, R. B.; Taconis, K. W. (1965). "A Realization of a London-Clarke-Mendoza Type Refrigerator". Low Temperature Physics LT9. p. 1253. (Tmin~220mK)
B.S. Neganov and co-workers in Dubna and H.E. Hall and co-workers in Manchester went below
100 mK (1966). Dubna rapidly reached 25 mK.
The principles and methods of dilution refrigeration have been substantially developed by J. Wheatley et al. at La Jolla.
Modern « wet » refrigerators are based on the Grenoble design by Frossati and coworkers.
The development of sintered silver heat exchangers Tmin led to Tmin~2 mK.
« Dry » refrigerators were developed
by K. Uhlig et al. on GM coolers (1993),
and on Pulse-tube coolers by K. Uhlig et al. (2002) and H. Godfrin (1999 to 2003, date of first commercial unit, delivered by l’ Air Liquide)
Lancaster has the present record of low temperatures, 1.75 mK

## Slide 5: Phase diagram of helium mixtures
By Mets501 - Own work, CC BY-SA 3.0, https://commons.wikimedia.org/w/index.php?curid=25250292

## Slide 6: Phase separation upon cooling

## Slide 7: Schematic diagram of a standard (“wet”) dilution refrigerator.
https://en.wikipedia.org/wiki/Dilution_refrigerator

## Slide 8: Low temperature part
https://en.wikipedia.org/wiki/Dilution_refrigerator

## Slide 9
https://en.wikipedia.org/wiki/Dilution_refrigerator

## Slide 10
Continuously Operating 4He Evaporation Refrigerator
L. E. DeLong, O. G. Symko and J. C. Wheatley, Rev. Sci. Inst. 42, 147 (1971)

## Slide 11: Enthalpy diagram of 4He

## Slide 12: Enthalpy diagram of 3He
Incoming 3He at P=0.5 bar and 4.2 K has a enthalpy of 99 J/mol
3He leaves the still at 0.7 K and P~0. It’s enthalpy is 35 J/mol.
The deficit is huge: 64 J/mol
Solutions: 1K pot or condensation at high pressure

## Slide 13: Global enthalpy balance of DR (only pure 3He flows in and out the sub-system)
Without 1K pot: negative balance,
External cooling needed.
With a 1K pot at 1.2 K: positive balance,
external heating (usually at the still…)
needed.
The cooling power provided by the 1 K pot to change the 3He from gas at 0.5 bar and 4.2 K to liquid at 0.5 bar and 1.2 K, is

## Slide 14: Condensing the 3He at ~3bars
J. Kraus, Cryogenics 17, 173 (1977)
A.T.A.M. de Waele, A.B. Reekers, H.M. Gijsman, Cryogenics 17, 175 (1977)
At 4.2 K the minimum enthalpy (vs. Pressure) is Hm ≈ 42 J/mol near 3 bars,
just above the enthalpy of the vapor at 0.7 K (35 J/mol).
> steady state is possible using a heat exchanger to precool the incoming 3He !!!

## Slide 15: Dilute side

## Slide 16: The Mixing Chamber: cooling is due to expansion

## Slide 17: Osmotic enthalpy

## Slide 19: Osmotic enthalpy of 3He mixtures
- Lines of constant osmotic pressure
(solid lines, values in Pa)
- Lines of constant osmotic enthalpy H3
(dotted lines values in J/mol);
and the saturation concentration of the
dilute phase xs.
The scale on the right gives the temperature in K.

## Slide 20: Cooling power at the MC

## Slide 21: Limiting temperature of DR